## Slide 1
Characterization and automated alignment detection of an additively manufactured Z-frame marker to process signals for robotic control in interventional MRI
Parisa Parsanejad
Parisa.parsanejad@yahoo.com
Medical Systems Engineering
Supervisors: Prof. Dr. Michael Friebe
M. Sc. Robert Odenbach
07.06. 2018

## Slide 2
Contents
                Introduction
                Processing methods
                Results
                Conclusion and outlook
                GUI live demonstration

## Slide 3
3D-printed Z-frame marker
Needle alignment to the image plane
A full six degree-of-freedom estimate of a device’s pose in
MRI coordinates using a single image slice of the frame
Notes: Z-frame is a passive tracking marker
control and tracking of the position and orientation of the instrument
designed with internal cylindrical tunnels, which are aligned in vertical and 45°-orientation at three sides (= Z-shape).
A single image slice will show seven ellipses, the centroids are used for localizing the frame
The benefit from the Z-shaped marker design is that the full 6-DOF (DOF = degree of freedom) pose of the device can be extracted from the relationship of the dots that show up on the image.
Our new marker concept uses the same design with the Z-shaped tunnels, but it is fully additively manufactured.

## Slide 4
3D-printed Z-frame marker
Internal Z-shaped tunnel structure without any entrance from the outside
New features:
            View correction triangle
Center pointing triangles
Fig .1: Z-RP-MR-marker with the distinct shaped internal tunnels in the CAD-view
Fig .2: Top view of the marker
Z= +10 mm
Fig .3: Bottom view of the marker
Z= -10 mm
P1
P1
P6
P5
P4
P3
P2
P7
P6
P5
P4
P3
P2
P7
Notes:             Clear identification of the correct view of the marker
Z-frame is a passive tracking marker
control and tracking of the position and orientation of the instrument
designed with internal cylindrical tunnels, which are aligned in vertical and 45°-orientation at three sides (= Z-shape).
A single image slice will show seven ellipses, the centroids are used for localizing the frame
The benefit from the Z-shaped marker design is that the full 6-DOF (DOF = degree of freedom) pose of the device can be extracted from the relationship of the dots that show up on the image.
Our new marker concept uses the same design with the Z-shaped tunnels, but it is fully additively manufactured.

## Slide 5
                Configuration of robot-assisted iMRI
Robotic
Device with
Attached
Z-frame Marker
Navigation
Workstation
Command
Status
Control Room
Scanner Room
MRI Host Workstation
Robot Controller
MRI Scanner
MRI Hardware

## Slide 6
Navigation workstation: Real-time localization algorithm
Real time MR imaging
Robot movement
to the goal position
Marker alignment calculation
Rigid transformation matrix
Angle calculation of image plane & frame
Instrument trajectory estimation
Robotic device motion calculation
Image pre-processing
stage
Localization and registration stage
Image filtering
and
segmentation
Center detection of the fiducials ellipses
Distance measurements between centroids
Validity test and calibration
Validity test and calibration

## Slide 7
Hardware setup for real-time image capturing
HD- video recorder
HDMI and USB cables
Converter
MRI setup with
capture device

## Slide 8
Automatic detection and segmentation of the Z-frame structure
Z-frame detection using faster R-CNN deep learning
                Input layer
                Middle layers
                Convolution2dLayer, reluLayer (rectified linear units), maxPooling2dLayer
                Final layers
                fullyConnectedLayer, softmaxLayer
Notes: Start with the imageInputLayer, which defines the type and size of the input layer. For detection tasks, the CNN needs to analyze smaller sections of the image, so the input size must be similar in size to the smallest object in the data set.
This input size is a balance between processing time and the amount of spatial detail the CNN needs to resolve.
Next, define the middle layers of the network. The middle layers are made up of repeated blocks of convolution2dLayer, reluLayer (rectified linear units), and maxPooling2dLayer layers. These layers form the core building blocks of convolutional neural networks.

## Slide 9
Image pre-processing
Noise
removal
Binary
masking
Image
binarization
Contrast
enhancement
Notes: The main drawback of the linear filtering is that the details in the original image will be destroyed during reducing the noise.
Anisotropic diffusion filtering proposed by Perona and Malic [reference] is a technique which reduces the image noise but preserves details and even enhances edges.

## Slide 10
Image pre-processing
Noise
removal
                2D Gaussian filter:
                Localization error = 0.7327
                Anisotropic diffusion
                Localization error= 0.6114
                Median Filter
                Localization error= 0.6114
Localization error:
the Euclidean distance between the predicted marker location, computed from the z-frame registration, and its true position as measured in MR images:
Z-frame registration:
Notes: The main drawback of the linear filtering is that the details in the original image will be destroyed during reducing the noise.
Anisotropic diffusion filtering proposed by Perona and Malic [reference] is a technique which reduces the image noise but preserves details and even enhances edges.
Localization error: Localization error is defined as the Euclidean distance between the predicted marker location, computed from the z-frame registration, and its true position as measured in the high-resolution MR images.

## Slide 11
Image filtering steps
Noise
removal
Binary
masking
Image
binarization
Contrast
enhancement
                2D Gaussian filter:
                Localization error = 0.7327
                Anisotropic diffusion
                Localization error= 0.6114
                Median Filter
                Localization error= 0.6114
Localization error:
the Euclidean distance between the predicted marker location, computed from the z-frame registration, and its true position as measured in MR images:
D=Pf-P
LE=norm(D)
Z-frame registration:
Notes: The main drawback of the linear filtering is that the details in the original image will be destroyed during reducing the noise.
Anisotropic diffusion filtering proposed by Perona and Malic [reference] is a technique which reduces the image noise but preserves details and even enhances edges.

## Slide 12
Image filtering steps
Noise
removal
Binary
masking
Image
binarization
Contrast
enhancement
                2D Gaussian filter:
                Localization error = 0.7327
                Anisotropic diffusion
                Localization error= 0.6114
                Median Filter
                Localization error= 0.6114
                Imadjust MATLAB
image= imadjust(image,[0.4 .7],[0 1])
Notes: The main drawback of the linear filtering is that the details in the original image will be destroyed during reducing the noise.
Anisotropic diffusion filtering proposed by Perona and Malic [reference] is a technique which reduces the image noise but preserves details and even enhances edges.

## Slide 13
Image filtering steps
Noise
removal
Binary
masking
Image
binarization
Contrast
enhancement
                2D Gaussian filter:
                Localization error = 0.7327
                Anisotropic diffusion
                Localization error= 0.6114
                Median Filter
                Localization error= 0.6114
                Imadjust MATLAB
image= imadjust(image,[0.4 .7],[0 1])
                Im2bw MATLAB
Notes: The main drawback of the linear filtering is that the details in the original image will be destroyed during reducing the noise.
Anisotropic diffusion filtering proposed by Perona and Malic [reference] is a technique which reduces the image noise but preserves details and even enhances edges.

## Slide 14
Image filtering steps
Noise
removal
Binary
masking
Image
binarization
Contrast
enhancement
                2D Gaussian filter:
                Localization error = 0.7327
                Anisotropic diffusion
                Localization error= 0.6114
                Median Filter
                Localization error= 0.6114
                Imadjust MATLAB
image= imadjust(image,[0.4 .7],[0 1])
                Im2bw MATLAB
                bwareaopen MATLAB
                Removal of small objects from the image
Notes: The main drawback of the linear filtering is that the details in the original image will be destroyed during reducing the noise.
Anisotropic diffusion filtering proposed by Perona and Malic [reference] is a technique which reduces the image noise but preserves details and even enhances edges.

## Slide 15
Image filtering steps
Noise
removal
Binary
masking
Image
binarization
Contrast
enhancement
                Imadjust MATLAB
image= imadjust(image,[0.4 .7],[0 1])
                Im2bw MATLAB
                bwareaopen MATLAB
                Removal of small objects from the image
Notes: The main drawback of the linear filtering is that the details in the original image will be destroyed during reducing the noise.
Anisotropic diffusion filtering proposed by Perona and Malic [reference] is a technique which reduces the image noise but preserves details and even enhances edges.

## Slide 16
Calibration:
Automatic rotation of the fiducial points
P1
P1
P1
P2
P3
P4
P5
P6
P7
P1
P2
P3
P4
P5
P6
P7

## Slide 17
Automatic rotation of the fiducial points
*
*
*
*
*
*
*
*
*
*
*
*
*
*
*
*
*
*
*
*
*
P1
P1
P1
P1
P7
P7
P7
P7
                Calculation of the Euclidean distance between each point and its neighbors
                Determination of the position of the P1 based on its neighbor's points
                Estimation of the rotation angle
                Theta= alpha-45+D
*
*
*
*
*
*
*
*
*
*
*
*
*
*
*
*
*
*
*
*
*
*
*
*
*
*
*
*
*
*
*
*
*
*
*
D=180
D=270
D=90
D=0
x
x
x
x
Y
Y
Y
Y

## Slide 18
Marker alignment calculation
[p2f p4f p6f]
= -20.0000 -1.1483 20.0000
-0.5648 20.0000 -7.4107
0.5648 1.1483 -7.4107
Rotation around x axis: 10 degree
Rotation around y axis: 10 degree
Notes: The ordered set of fiducial point coordinates P’ are then applied to compute the 6-DOF pose of the Z-frame with respect to the image plane. Finally, the computed frame position and orientation is used to compute the required motion of the robotic device to reach the target pose.

## Slide 19
Conclusion and Outlook
                conclusion
                Phantom targeting test with Z-RP- frame and an MR-compatible robotic device
                Improving the real-time image acquisition setup
                Test setup

## Slide 20
Thank you for your attention!
Parisa Parsanejad
Parisa.parsanejad@ovgu.de
INKA - Catheter Technologies
Institute for Medical Systems
Faculty of Electrical Engineering and Information Technologies
Otto-von-Guericke-University Magdeburg
Rötgerstr. 9, D-39104 Magdeburg
www.inka-md.de
This work is funded by the BMBF (03IPT7100X)

## Slide 21
Backup

## Slide 22
Evaluate detector using test set
Notes: The precision/recall (PR) curve highlights how precise a detector is at varying levels of recall. Ideally, the precision would be 1 at all recall levels.
The use of additional layers in the network can help improve the average precision, but might require additional training data and longer training time.

## Slide 23
Calibration:
Automatic rotation of the fiducial points
P1
P1
P1
P1
P1
P1