## Slide 1
Characterization and automated alignment detection of an additively manufactured Z-frame marker to process signals for robotic control in interventional MRI
Parisa Parsanejad
Medical Systems Engineering
Supervisors:
Prof. Dr. Michael Friebe
M. Sc. Robert Odenbach
Final Presentation
24.10.2018

## Slide 2
Contents
                Introduction
                MRI-guided intervention
                Additively manufactured Z-frame
                Z-frame marker design and characterization
                Methodology
                Alignment workflow
                Navigation workstation
                Graphical user interface
                Testing and results
                Gold standard CAD analysis
                Z-frame marker positioning
                Phantom targeting test
                Conclusion and outlook

## Slide 3
MRI-guided interventions
Application: Biopsy, ablation, neurosurgical interventions
Methods: Freehand approach
MR-compatible mechanical or robotic device
Notes: The ordered set of fiducial point coordinates P’ are then applied to compute the 6-DOF pose of the Z-frame with respect to the image plane. Finally, the computed frame position and orientation is used to compute the required motion of the robotic device to reach the target pose.

## Slide 4
Passive tracking Z-frame marker
Internal cylindrical tunnel structure with 45°-orientation at three sides
Full six degree-of-freedom estimation of a device’s pose in MRI coordinates using a single image slice of the frame
P2
P3
P1
P4
P6
P5
P7
P3
P1
P5
P7
Notes: Z-frame is a passive tracking marker
control and tracking of the position and orientation of the instrument
designed with internal cylindrical tunnels, which are aligned in vertical and 45°-orientation at three sides (= Z-shape).
A single image slice will show seven ellipses, the centroids are used for localizing the frame
The benefit from the Z-shaped marker design is that the full 6-DOF (DOF = degree of freedom) pose of the device can be extracted from the relationship of the dots that show up on the image.
Our new marker concept uses the same design with the Z-shaped tunnels, but it is fully additively manufactured.

## Slide 5
Passive tracking Z-frame marker
Internal cylindrical tunnel structure with 45°-orientation at three sides
Full six degree-of-freedom estimation of a device’s pose in MRI coordinates using a single image slice of the frame
P2
P3
P1
P4
P6
P5
P7
P3
P1
P5
P7
Notes: Z-frame is a passive tracking marker
control and tracking of the position and orientation of the instrument
designed with internal cylindrical tunnels, which are aligned in vertical and 45°-orientation at three sides (= Z-shape).
A single image slice will show seven ellipses, the centroids are used for localizing the frame
The benefit from the Z-shaped marker design is that the full 6-DOF (DOF = degree of freedom) pose of the device can be extracted from the relationship of the dots that show up on the image.
Our new marker concept uses the same design with the Z-shaped tunnels, but it is fully additively manufactured.

## Slide 6
Z-frame physical structures
            New features:
            Square-shaped tunnel structure
            Z-axis indication triangle
            Center pointing triangles
Top view of the marker
Z= +10 mm
Bottom view of the marker
Z= -10 mm
P1
P6
P5
P4
P3
P2
P7
P1
P6
P5
P4
P3
P2
P7
P4
P7
P6
P5
P3
P2
P1
Notes:             Clear identification of the correct view of the marker
Z-frame is a passive tracking marker
control and tracking of the position and orientation of the instrument
designed with internal cylindrical tunnels, which are aligned in vertical and 45°-orientation at three sides (= Z-shape).
A single image slice will show seven ellipses, the centroids are used for localizing the frame
The benefit from the Z-shaped marker design is that the full 6-DOF (DOF = degree of freedom) pose of the device can be extracted from the relationship of the dots that show up on the image.
Our new marker concept uses the same design with the Z-shaped tunnels, but it is fully additively manufactured.
            Center pointing triangles: Indirect detection of the center of the marker
            Z-axis indication triangle: Clear identification of the correct view orientation of the marker
Z-RP-MR-marker with the distinct shaped internal tunnels in the CAD-view

## Slide 7
                Workflow of the nearly real-time alignment detection
Nearly real-time
image acquisition
Image processing
Robotic/mechanical device translation
Micro positioning device image
Rotational Alignment detection
Alignment controller
processing
Translational Alignment detection
Alignment controller
processing
Robotic/mechanical device rotation

## Slide 8
                Workflow of the nearly real-time alignment detection
Nearly real-time
image acquisition
Image processing
Robotic/
mechanical device
translation
Micro positioning device image
Rotational alignment detection
Alignment controller
processing
Translational alignment detection
Alignment controller
processing
Robotic/
mechanical device
rotation
MRI host
workstation

## Slide 9
Hardware setup for real-time image capturing
HD- video recorder
AV-converter
MRI
monitor
Processing laptop

## Slide 10
Real time image capturing
Robot movement to the target trajectory
Marker alignment calculation
Rigid transformation matrix
Angle calculation of image plane & frame
Instrument trajectory estimation
Image pre-processing
stage
Localization and registration stage
Image filtering
and
segmentation
Center detection of the fiducials ellipses
Distance measurements between centroids
Validity test and calibration
Marker alignment calculation
Navigation workstation: Real-time localization algorithm
Robotic device motion calculation
Rigid transformation matrix
Angle calculation of image plane & frame
Instrument trajectory estimation

## Slide 11
                Graphical user interface
Automated alignment detection GUI

## Slide 12
                Graphical user interface
Automated alignment detection GUI
3. Z-frame Segmentation

## Slide 13
Z-frame detection using faster R-CNN (Regions with Convolutional Neural Networks)
          CNN layer:
Input layer
                Middle layers: convolution2dLayer, reluLayer (rectified linear units), maxPooling2dLayer
                Final layers: fullyConnectedLayer, softmaxLayer
Precision
Recall
Average Precision = 0.35
3. Z-frame Segmentation
Automatic button
Notes: Start with the imageInputLayer, which defines the type and size of the input layer. For detection tasks, the CNN needs to analyze smaller sections of the image, so the input size must be similar in size to the smallest object in the data set.
This input size is a balance between processing time and the amount of spatial detail the CNN needs to resolve.
Next, define the middle layers of the network. The middle layers are made up of repeated blocks of convolution2dLayer, reluLayer (rectified linear units), and maxPooling2dLayer layers. These layers form the core building blocks of convolutional neural networks.

## Slide 14
                Graphical user interface
                4. Image Pre-processing
Automated alignment detection GUI

## Slide 15
4. Image Pre-processing
Noise
removal
Binary
masking
Image
binarization
Contrast
enhancement
                2D Gaussian filter
                Wavelet filter
                Anisotropic diffusion
Image intensity values adjustment
Image= imadjust(Image,[0.5 .9],[0 1])
                Image thresholding
Image= im2bw(Image,0.5);
                Removal of small objects from the image
                mask = bwareaopen(yy,50);
                mask = imclearborder(mask);
% North, South, East and West differences
deltaN = diffl(1:rows,2:cols+1) - diff;
deltaS = diffl(3:rows+2,2:cols+1) - diff;
deltaE = diffl(2:rows+1,3:cols+2) - diff;
deltaW = diffl(2:rows+1,1:cols) - diff;
% Conduction
if option == 1
cN = exp(-(deltaN/kappa).^2);
cS = exp(-(deltaS/kappa).^2);
cE = exp(-(deltaE/kappa).^2);
cW = exp(-(deltaW/kappa).^2);
elseif option == 2
cN = 1./(1 + (deltaN/kappa).^2);
cS = 1./(1 + (deltaS/kappa).^2);
cE = 1./(1 + (deltaE/kappa).^2);
cW = 1./(1 + (deltaW/kappa).^2);
end
diff = diff + lambda*(cN.*deltaN + cS.*deltaS + cE.*deltaE + cW.*deltaW);
end
Notes: The main drawback of the linear filtering is that the details in the original image will be destroyed during reducing the noise.
Anisotropic diffusion filtering proposed by Perona and Malic [reference] is a technique which reduces the image noise but preserves details and even enhances edges.

## Slide 16
                Graphical user interface
                5. Center Detection
Automated alignment detection GUI
s = regionprops(mask,'centroid');
centroids = cat(1, s.Centroid);
centroids = round(centroids);
figure
imshow(Image)
hold on
plot(centroids(:,1),centroids(:,2), 'b*')
hold off
[m n]=size(centroids)

## Slide 17
                Graphical user interface
                6. Calibration
Automated alignment detection GUI
                R = [cos(theta) -sin(theta); sin(theta) cos(theta)];
                x_intersect = fzero(@(x) polyval(p1-p2,x),3)
                y_intersect = polyval(p1,x_intersect)
                line(x1,y1);
                line(x2,y2);
                % transfer the points into new coordinate system
                P(:,1)=P(:,1)-x_intersect
                P(:,2)=P(:,2)-y_intersect
                % Convert Pixel to mm
                [ps_camera_x,ps_camera_y] = camera_calibration(P)
                P(:,1)=P(:,1)*ps_camera_x
                P(:,2)=P(:,2)*ps_camera_y

## Slide 18
                Graphical user interface
                7. Alignment detection
Automated alignment detection GUI

## Slide 19
                Graphical user interface
                8. Robotic Rotation Adjustment
Automated alignment detection GUI

## Slide 20
Testing and analysis
Sources of error
Alignment detection algorithm
Capture device resolution
MR-imaging artifacts
Z-frame printing artifact (air bubbles)
positioning frame accuracy

## Slide 21
Testing and analysis
Scenarios
                Scenario 1: Gold standard CAD analysis
                Scenario 2: Z-frame marker positioning
                using an MR-compatible rotary frame
                Scenario 3: Phantom targeting test
                Rotation of the Z-frame around the x-axis
                Rotation of the Z-frame around the y-axis
                Rotation of the Z-frame around both the x- and y-axis with the same orientation
                Rotation of the Z-frame around both the x- and y-axis with the opposite orientation

## Slide 22
Result analysis: Scenarios CAD ………. and 2

## Slide 23
Regression analysis
Rotation of the Z-frame around (degree) | R² for the gold standard rotation around the x-axis | R² for the gold standard rotation around the y-axis | R² for the MRI rotation around the x-axis | R² for the MRI rotation around y-axis
X-axis | 0.9997 | - | 0.9981 | -
Y-axis | - | 0.9999 | - | 0.9991
X- and Y-axes (same direction) | 0.999 | 0.9974 | 0.9998 | 0.9989
X- and Y-axes (opposite direction) | 0.9998 | 0.9958 | 0.9997 | 0.9975
Average R² | 0.9995 | 0.9977 | 0.9992 | 0.9985
Total R² | 0.9986 |  | 0.9988 | 

## Slide 24
Root mean squared error analysis
Rotation of the Z-frame around (degree) | RMSE for the gold standard rotation around the x-axis | RMSE for the gold standard rotation around the y-axis | RMSE for the MRI rotation around the x-axis | RMSE for the MRI rotation around y-axis
X-axis | 0.18 | - | 0.49 | -
Y-axis | - | 0.13 | - | 0.51
X- and Y-axes (same direction) | 0.32 | 0.66 | 0.18 | 0.38
X- and Y-axes (opposite direction) | 0.26 | 0.21 | 1.0 | 0.47
Average RMSE | 0.25 | 0.33 | 0.56 | 0.45
Total RMSE | 0.29 |  | 0.50 | 
Alignment detection algorithm
Alignment detection algorithm
Capture device resolution
MR-imaging,
Z-frame printing artifact
positioning frame

## Slide 25
Scenario 3: Phantom targeting test
Rotation around x-axis= -6.04⁰
Rotation around y-axis= -3.81⁰
Rotation around x-axis= -0.52⁰
Rotation around y-axis= -0.62⁰

## Slide 26
Phantom targeting test
40mm

## Slide 27
Phantom targeting test

## Slide 28
Conclusion and Outlook
                The proposed registration method can contribute to a faster and more precise guiding of the MRI-based interventions Image acquisition time: 20 sec
                Rotation processing time : 15 -20 sec
                Rotation adjustment time: 2 min
                Image acquisition time: 20 sec
                Translation processing: 15 sec
                Translation adjustment time : 2 min
                RMSE : 0.5 degree
                No need for an additional interface for connecting the MRI scanner and the navigation software
                Automatic computation of the translational alignment of the Z-frame marker
Improvement in the automatic segmentation of the Z-frame using a machine learning technique
For the image acquisition setup, a new capture device with a higher quality
Better converter ( the tested ones)

## Slide 29
Conclusion and Outlook
                Automatic computation of the translational alignment of the Z-frame marker
Improvement in the automatic segmentation of the Z-frame using a machine learning technique
For the image acquisition setup, a new capture device with a higher quality
Better converter ( the tested ones)
        WORKFLOW
        IMRI procedures (FINISH) (VISION FULLFILL)

## Slide 30
References
Cepek, Jeremy, Blaine A. Chronik, and Aaron Fenster. "Errors in device localization in MRI using Z-Frames." International Conference on Medical Image Computing and Computer-Assisted Intervention. Springer, Berlin, Heidelberg, 2013.
2. Su, Hao, et al. "A concentric tube continuum robot with piezoelectric actuation for MRI-guided closed-loop targeting." Annals of biomedical engineering 44.10 (2016): 2863-2873.
3. DiMaio, Simon P., et al. "Dynamic MRI scan plane control for passive tracking of instruments and devices." International Conference on Medical Image Computing and Computer-Assisted Intervention. Springer, Berlin, Heidelberg, 2007.
Tokuda, Junichi, et al. "Configurable automatic detection and registration of fiducial frames for device-to-image registration in MRI-guided prostate interventions." International Conference on Medical Image Computing and Computer-Assisted Intervention. Springer, Berlin, Heidelberg, 2013.
                Cepek, J., Chronik, B. A., & Fenster, A. (2014). The effects of magnetic field distortion on the accuracy of passive device localization frames in MR imaging. Medical physics, 41(5).
                Odenbach, R.; Grundmann, M.; Boese, A.; Friebe, M.: Generatively manufactured, passive and safe MRI-marker. In 11th Interventional MRI Symposium, Baltimore, 2016
                Odenbach, R.; Chen, Y.; Sengupta, S.; Webster, R.; Barth, E.; Friebe, M.: 3D-printed Z-frame marker for MRI-guided interventions. In 16. Jahrestagung der Deutschen Gesellschaft fuer Computer- und Roboterassistierte Chirurgie, Hannover, 2017
                Mattern, H.; Odenbach, R.; Parsanejad, P.; and Friebe, M.: 3D-printed MRI marker for personalized interventional applications through T1 and T2 relaxation time matching. In CARS 2018 Computer Assisted Radiology and Surgery, Berlin, Germany, 2018

## Slide 31
Thank you for your attention!
Parisa Parsanejad
Parisa.parsanejad@ovgu.de
INKA - Catheter Technologies
Institute for Medical Systems
Faculty of Electrical Engineering and Information Technologies
Otto-von-Guericke-University Magdeburg
Rötgerstr. 9, D-39104 Magdeburg
www.inka-md.de
This work is funded by the BMBF (03IPT7100X)

## Slide 32
Automated alignment detection algorithm

## Slide 33
Backup

## Slide 34
Evaluate detector using test set
Notes: The precision/recall (PR) curve highlights how precise a detector is at varying levels of recall. Ideally, the precision would be 1 at all recall levels.
The use of additional layers in the network can help improve the average precision, but might require additional training data and longer training time.

## Slide 35
Calibration:
Automatic rotation of the fiducial points
P1
P1
P1
P1
P1
P1

## Slide 36
Image filtering steps
Noise
removal
Binary
masking
Image
binarization
Contrast
enhancement
                2D Gaussian filter:
                Localization error = 0.7327
                Anisotropic diffusion
                Localization error= 0.6114
                Median Filter
                Localization error= 0.6114
Localization error:
the Euclidean distance between the predicted marker location, computed from the z-frame registration, and its true position as measured in MR images:
D=Pf-P
LE=norm(D)
Z-frame registration:
Notes: The main drawback of the linear filtering is that the details in the original image will be destroyed during reducing the noise.
Anisotropic diffusion filtering proposed by Perona and Malic [reference] is a technique which reduces the image noise but preserves details and even enhances edges.

## Slide 37
Marker alignment calculation
[p2f p4f p6f]
= -20.0000 -1.1483 20.0000
-0.5648 20.0000 -7.4107
0.5648 1.1483 -7.4107
Rotation around x axis: 10 degree
Rotation around y axis: 10 degree
Fig .16: Surface of the three inner points
Notes: The ordered set of fiducial point coordinates P’ are then applied to compute the 6-DOF pose of the Z-frame with respect to the image plane. Finally, the computed frame position and orientation is used to compute the required motion of the robotic device to reach the target pose.

## Slide 38
                Z-frame marker characterization
Result T1 and T2

## Slide 39
Image rotation estimation using automated feature matching

## Slide 40
Calibration:
Automatic rotation of the fiducial points
P1
P1
P1
P2
P3
P4
P5
P6
P7
P1
P2
P3
P4
P5
P6
P7
Notes: Talk about different approach for rotation like feature extraction SURF and SIFT and MSER but they didnt work for all the images because of the differeces in circle shapes

## Slide 41
Automatic rotation estimation using centroids coordinate system
*
*
*
*
*
*
*
*
*
*
*
*
*
*
*
*
*
*
*
*
*
P1
P1
P1
P1
P7
P7
P7
P7
*
*
*
*
*
*
*
*
*
*
*
*
*
*
*
*
*
*
*
*
*
*
*
*
*
*
*
*
*
*
*
*
*
*
*
D=180
D=270
D=90
D=0
x
x
x
x
Y
Y
Y
Y
P1
P1
P1
P1
P2
P2
P2
P2
P3
P4
P4
P4
P4
P3
P3
P3
P5
P5
P5
P5
P6
P6
P6
P6
P7
P7
P7
P7
                Calculation of the Euclidean distance between each point and its neighbors
                Determination of the position of the P1 based on its neighbor's points
                Estimation of the rotation angle
                𝜃= 𝛼-45+D
Fig .15: Four different possible rotation of the Z-marker

## Slide 42
Testing

## Slide 43
STEP 1: Z-frame marker positioning using an MR-compatible rotary frame
                Rotation of the Z-frame around the x-axis
                Rotation of the Z-frame around the y-axis
                Rotation of the Z-frame around both the x- and y-axis with the same orientation
                Rotation of the Z-frame around both the x- and y-axis with the opposite orientation
                Rotation of the Z-frame around both the x- and y-axis with the random angles